## Slide 1
Digitizers for Physics Applications
RT2014 Nara – May 27th 2014
Carlo Tintori
ISO 9001:2008
CERT. N. 9105.CAEN

## Slide 2: CAEN Digitizers Highlights
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
VME, NIM, Desktop form factors
Up to 64 channels in a single board
8, 10, 12 and 14 bit; up to 5GS/s
Digital MCA with integrated HV and LV
FPGA firmware for Digital Pulse Processing:
  Pulse Height Analysis
  Pulse Shape Discrimination
  Charge Integration
  CFD and Time stamping
Readout:
  VME up to 160MB/s, shared bus
  Optical Link up to 100 MB/s point to point (PCI/PCIe)
  USB 2.0 up to 30MB/s
Memory buffer: up to 10MB/ch
Multi-board synchronization: clock, synch and trigger distribution
Programmable digital I/Os
On-line coincidence and majority
Software for Windows and Linux
Notes: 2

## Slide 3: Digitizers Table
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
MODEL(1) | Form Factor | # channels | Sampling Frequency (MS/s) | # Bits | Input Dynamic Range (Vpp) | Bandwidth (MHz) | Memory (Msample/ch) Record Lenght |  | DPP firmware(4)
 |  |  |  |  |  |  | small | big | 
x724 | VME | 8 | 100 | 14 | 0.5 - 2.25 - 10 | 40 | 0.5 500s | 4 4ms | PHA
 | Desktop/NIM | 4 |  |  |  |  |  |  | 
x720 | VME | 8 | 250 | 12 | 2 | 125 | 1.25 5ms | 10 40ms | CI, PSD
 | Desktop/NIM | 4 |  |  |  |  |  |  | 
x730 | VME | 16 | 500 | 14 | 0.5 and 2 (5) | 250 | 1.25 250s | 10 2ms | PHA, PSD(6)
 | Desktop/NIM | 8 |  |  |  |  |  |  | 
x721 | VME | 8 | 500 | 8 | 2 | 250 | 2 2ms |  | -
x731 | VME | 8/4 | 500/1000 | 8 | 2 | 250/500 | 2/4 2ms |  | -
 | Desktop/NIM | 4/2 |  |  |  |  |  |  | 
x751 | VME | 8 | 1000/2000 | 10 | 1 | 500 | 1.8/3.6 1.8ms | 14.4/28.8 14.4ms | PSD, ZLE
 | Desktop/NIM | 4 |  |  |  |  |  |  | 
x761 | VME | 2 | 4000 | 10 | 1 | 1000 | 7.2 1.8ms | 57.6 14.4ms | -
 | Desktop/NIM | 1 |  |  |  |  |  |  | 
x740 | VME | 64 | 62.5 | 12 | 2 - 10 | 30 | 0.19 3ms | 1.5 24ms | -
 | Desktop/NIM | 32 |  |  |  |  |  |  | 
x742 | VME | 32+2 | 5000(2) | 12 | 1 | 600 | 0.128(3) 200ns |  | -
 | Desktop/NIM | 16+1 |  |  |  |  |  |  | 
x743 | VME | 16 | 3200(2) | 12 | 2.5 | 500 | 0.003(7) 640ns |  | CM
 | Desktop/NIM | 8 |  |  |  |  |  |  | 
(1) The x in the model name is V1 for VME, VX1 for VME64X, DT5 for Desktop and N6 for NIM
(2) Sampling frequency of the analog memory (switched capacitor array); A/D conversion takes place at lower speed (dead-time)
(3) The memory size for the x742 is 128 events of 1024 samples each. Record length can be 200 ns, 500 ns or 1 s depending on the sampling frequency
(4) DPP-PHA: Pulse Height analysis (Trapezoidal Filters), DPP-CI and CM: Charge Integration (digital QDC); DPP-PSD: n/γ Discrimination (double gate charge)
(5) Input dynamic range with 2 options software selectable
(6) PSD available from Q4-2013, PHA available from Q4-2014
(7) The memory size for the x743 is 3 events of 1024 samples each
Notes: 3

## Slide 4: Readout Modes
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
CONET2 (1.25 Gb/s Optical Link):
  Up to ~100MB/s per link
  Daisy chainable (up to 8 boards)
  A2818 PCI (1 link), A3818 PCIe (1, 2 or 4 links)
VME64X
  MBLT (~70MB), 2eVME and 2eSST (~150MB) data readout
  Chained Block Transfer
  Backplane shared between boards (shared bandwidth)
USB 2.0
  ~30MB/s max bandwidth
  No external hardware required
  Easy to use (e.g. portable systems, laptops, etc…)
Notes: 4

## Slide 5: CONET2 readout example: XMASS
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
64 V1751 modules in 4 VME crates
512 channels (10 bit @ 1GHz)
4 A3818s 4 link PCIe cards
16 parallel CONET2 links
4 digitizers daisy chained
Readout Bandwidth = ~2 MB/s/ch
Total Bandwidth = ~ 1GB/s
Notes: 5

## Slide 6
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Designed for System scalability (Modular Electronics)
Multi Board Synchronization
On line Digital Pulse Processing (DPP):
  Individual pulse triggering
  Time stamp with LED of CFD discriminators + interpolation
  Baseline subtraction
  Energy calculation (pulse height or charge)
  Pulse Shape Discrimination
  Waveform Readout (optional)
  Coincidence or trigger propagation between channels
Digitizers for Physics applications
On-line DPP allows for a significant throughput rate reduction; this is mandatory for most physics applications
Notes: 6

## Slide 7: Traditional acquisition chain
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Detector
pre
Discr
TDC
Delay
Amplif
Peak
Sensing
splitter
Scaler
Energy
Charge
Counts
Time
ADC
QDC
Traditional acquisition chains are made of a number of analog modules interconnected with cables
Coinc
Gate
A/D conversion at the end of the chain
Notes: 7

## Slide 8: Fully digital acquisition chain
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Detector
pre
FPGA
Waveform
Charge
Count
Time
ADC
Energy
Digitizer
Algorithms
The aim of the DPP is to make a “all in digital” version of analog modules such as Shaping Amplifiers, Discriminators, QDCs, Peak Sensing ADCs, TDCs, Scalers, Coincidence Units, etc.
Nowadays fast and high resolution flash ADCs allow designers to make acquisition systems in which the A to D conversion occurs as close as possible to the detector
Shape
Notes: 8

## Slide 9
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
One single board can do the job of several analog modules
A/D conversion as early as possible, data reduction as late as possible: preserve full information!
Correlation of different information (e.g. timing and energy available from the same data readout): multi-parametric analysis
Reduction in size, cabling, power consumption and cost per channel
High reliability and reproducibility
Flexibility: different digital algorithms can be designed and loaded at any time into the same hardware
Benefits of the digital approach
Notes: 9

## Slide 10
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
DIGITAL DETECTOR EMULATOR
Notes: 10

## Slide 11: What is a Digital Detector Emulator
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
NOT a pulse generator of recorded shapes
BUT a synthesizer of true random pulses compliant to programmable statistics for amplitude, starting time and shape
Notes: 11

## Slide 12: DDE operating mode
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
MCA
ENERGY SPECTRUM
TIME DISTRIBUTION
PULSE SHAPE
NOISE + BASELINE DRIFT
ENERGY SPECTRUM
EMULATED SIGNALS
DDE
Notes: 12

## Slide 13: DDE: a full featured instrument
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
PULSER
TAIL GENERATOR
ps PULSE GENERATOR
WAVEFORM GENERATOR
WAVEFORM RECORDER
Notes: 13

## Slide 14: DDE Highlights
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Two analog channels (125MHz, 14bit DAC) + 4 digital I/Os
Random or periodic Pulse Generator emulating a programmable energy spectrum and a poissonian statistic emission with pile-up
Programmable Pulse Shape; dual-shape option for pulse shape discrimination tests
Transistor Reset Preamplifier emulation
Noise emulation (1/f, baseline drift, white noise, interference)
Correlated event emulation (with given energy spectrum and delay in steps of 11 ps) mixed in an uncorrelated background
Arbitrary Waveform generator (sine, square, triangular, pulses, …)
Isotopes database
Software and drivers for Windows
Notes: 14

## Slide 15: DDE Application Example 1
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Real Setup
Emulated Setup
Files
Rate
reproduction at high rate emulating a strong source
signals recorded with a weak source (low rate)
Educational: MCA training to students without sources and detectors
Notes: 15

## Slide 16: DDE Application Example 2
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Neutron Source
g ray
neutron
SHAPE 1
SHAPE 2
20%
80%
shape1/shape2 ratio
SHAPE 2
SHAPE 1
Notes: 16

## Slide 17: DDE Application Example 3
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
CH1
CH2
22Na
Uncorrelated background
Correlated Events
CH1
CH2
Notes: 17

## Slide 18
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
PULSE HEIGHT ANALYSIS
Notes: 18

## Slide 19: DPP-PHA topics
Digital Multi-Channel Analyzer:
Replaces the analog chain with shaping amplifier + peak sensing ADC
Takes the output signal of the Charge Sensitive Preamplifier
Implemented in the 14 bit, 100MSps digitizers (x724 models) and DT5780 digital MCA (that includes also High and Low Voltage supply) and DT5781. Coming soon on x730 modules (14 bit, 500MSps).
Provides pulse height, time stamp and optionally raw waveforms
Pile-up rejection, Baseline restoration, ballistic deficit correction
Best suited for high resolution spectroscopy (HPGe and Si detectors) as well as low cost solutions such as NaI, CsI, etc…
Multiple channel systems for clover or segmented detectors, compton suppression, cosmic veto, low background spectroscopy, etc…

## Slide 20: DPP-PHA Block Diagram
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Trigger Threshold
Trigger & Timing
Filter
Event
Builder
Waveforms
Self-Trigger
Input
PUR
TimeStamp
Output
Data
+
-
Deci
mator
Energy
Filter
ZC
Arm
Counter
clk
Baseline
Peak
Energy
Memory
Buffers
sync
Coinc
Reject
Freeze
Notes: 20

## Slide 21: DPP-PHA signals
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Notes: 21

## Slide 22: Pile-up in the Trapezoidal Filter
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Case 1: T > TTR+TTF (2nd trapezoid starts on the falling edge of the 1st one). Both energies are good (no pile-up events)
Case 2: ~TPR < T < TTR+TTF (2nd trapezoid starts on the rising edge or flat top of the 1st one). Pulse height calculation is not possible, no energy information is available (pile-up events); still two time stamps.
Case 3: T < ~TPR (input pulses piling up on their rising edge). The TT filter doesn’t distinguish the double pulse condition. Only one event is recorded (energy sum). The Rise Time Discriminator might mitigate this unwanted effect.
Notes: 22

## Slide 23: Test Results with HPGe and DT5780
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Energy (KeV) | Centroid (KeV) | FWHM (KeV)
59.541 | 59.521 +- 0.004 | 0.95 +- 0.01
68.895 | 68.910 +- 0.010 | 1.04 +- 0.03
70.819 | 70.809 +- 0.006 | 0.98 +- 0.01
661.659 | 661.716 +- 0.008 | 1.37 +- 0.01
1173.240 | 1173.233 +- 0.006 | 1.68 +- 0.01
1332.508 | 1332.487 +- 0.008 | 1.77 +- 0.01
1460.822 | 1460.833 +- 0.033 | 1.76 +- 0.06
DPP parameter | value
Pole zero time constant | 47 s
# Samples for Baseline | 1024
Trapezoid Rise Time | 5.0 s
Trapezoid Flat Top | 2.0 s
Peaking Delay | 1.5 s
Baseline Holdoff | 0.1 s
Peak Holdoff | 20.0 s
Test Conditions | 
MCA | DT5780
Detector | Canberra coaxial HPGe Mod. 7229P
Preamplifier | model 2001
HV bias | 4.5 kV
Sources | 60Co, 137Cs, 241Am, 204Tl
Counting rate | 100 - 500 Hz
Measured Dead time | < 2%
Notes: 23

## Slide 24: Test Results with HPGe and DT5724 (I)
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Notes: 24

## Slide 25
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Test Results with HPGe and DT5724 (III)
Notes: 25

## Slide 26
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
MC2 Analyzer
Notes: 26

## Slide 27
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
CHARGE INTEGRATION &
PULSE SHAPE DISCRIMINATION
Notes: 27

## Slide 28: DPP-CI/PSD topics
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Digital implementation of the QDC + discriminator and gate generator
CI: single gate; PSD: dual gate (fast and slow component)
Pulse Shape Discrimination for n-γ separation PSD = (QLONG - QSHORT)/QLONG
Available for x720 (12 bit @ 250MS/s), x751 (10 bit @ 1GS/s) and x730 (14 bit @ 500MS/s) and DT5790 (2 channel 12 bit @ 250MS/s + 2 High Voltage + 2 Low Voltage for preamps)
Digital CFD with time stamp interpolation (in x730 only)
Self-gating integration; no delay line to fit the pulse within the gate
Pile-up rejection or gate re-triggering
PSD cut suppress events with above/below a programmable PSD threshold (e.g. suppress gammas in low neutron counting rate cases)
Typically used with scintillators + PMT or SiPM/MPPC
Notes: 28

## Slide 29: DPP-PSD Block Diagram
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Trigger Threshold
Discr
Delay
Baseline
Charge
Accumulator
Event
Builder
Short
Gate
Long
Gate
PSD
PSD Threshold
QL
QS
Waveform
Trigger
Input
PUR
Time Stamp
Clock
Counter
Data
+
-
Notes: 29

## Slide 30: -n Discrimination: test results (I)
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Detector: BC501A 5x2 inches,
PMT: Hamamatsu R1250
Board: DT5270 with DPP-PSD
Notes: 30

## Slide 31: -n Discrimination: test results (II)
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Notes: 31

## Slide 32: -n Discrimination: test results (III)
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Notes: 32

## Slide 33: -n Discrimination: Comparing boards
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Notes: 33

## Slide 34: -n Discrimination: bits and sampling rate
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Big improvement from 250 to 500 MS/s
10 bits are not sufficient; 14 are not significant; 12 is OK
Notes: 34

## Slide 35
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
TIMING ANALYSIS
Notes: 35

## Slide 36: Conventional TDCs vs Digitizers
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Conventional TDC boards:
  V1190: 128 channel, 100 ps Multi-Hit TDC
  V1290: 32 channel, 25 ps Multi-Hit TDC
  V775: 32 channel, 35 ps Start-Stop TDC
TDC in a digitizer can't compete in terms of density and cost, but there are cases where the implementation of a TDC in a digitizer is profitable:
  Applications that require an excellent timing resolution (5 ps) of the whole chain (CFD+TDC)
  Simultaneous acquisition of Timing and Energy: digitizers do both
  Walk correction with the energy
  Bursts of very close pulses (e.g. Free Electron Lasers): the digitizer can operate without dead time
  Direct connection from detector to digitizer (no discriminators!). Less cables, less distortion… eventually less cost!
external
CFD required!
Notes: 36

## Slide 37: Algorithms for the Time Measurements
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Digital CFD: SCFD(n) = A*SIN(n) - SIN(n-D), A=Attenuation, D=Delay
Interpolation between samples necessary to improve timing resolution beyond the granularity of the sampling period
Linear interpolation (segment between two samples) is simple and effective in most cases; can be implemented in the FPGA
timing resolution is function of the sampling frequency, pulse amplitude and rise-time (i.e. V/T)
Rule of thumb: 4-5 samples in the leading edge to get good and stable timing resolution
Notes: 37

## Slide 38
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Timing Resolution (ps RMS)
 | Leading Edge |  |  |  | 
Board Type | 0.8 ns | 1.6 ns | 2.5 ns | 5 ns | 10 ns
2GS/s 10bit (x751) | 14 | 9 | - | - | -
4GS/s 10bit (x761) | 6 | 6 | - | - | -
Switched Capacitor ADCs (x742/x743) | 5 | 5 | - | - | -
500MS/s 14 bit (x730) (*) | - | - | 126 | 30 | 10
500MS/s 14 bit (x730) (*) + input shaper | - | 12 | - | - | -
Analog CFD+TDC | 50 | 50 | 43 | - | -
START-STOP PULSES: 1 Vpp
(*) digital CFD + time stamp interpolation on-line in the x730 (no waveform readout) and soon in the x743 too.
Notes: 38

## Slide 39
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
SINLE BOARD for MULTI ANALYSIS
Notes: 39

## Slide 40: An example: Phoswitch+LaBr detectors
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Requirements:
Phoswitch detectors (e.g. NaI + LaBr3) require Pulse Shape Discrimination to separate events that interact in one or in the other crystal
Wide energy dynamic range (40KeV to 40MeV)
Precise Timing information
Multi-parametric analysis
PSD and Energy windowing
One single board (x730) directly connected to the detectors (either LaBr3 or Phoswitch) is able to manage the acquisition
Notes: 40

## Slide 41: Waveforms
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Phoswitch: Different Pulse Shape for interactions in LaBr3, NaI or scattered in both (Compton)
Notes: 41

## Slide 42: Energy Spectrum
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Energy Range up to 44MeV
Resolution in LaBr3: 19.2 KeV @ 662 KeV (2.9%)
8.6 KeV @ 80 KeV (10.7%)
Notes: 42

## Slide 43: Time of Flight Spectrum
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
3”x3” LaBr3 to Phoswitch Time Of Flight
Resolution = down to ~500 ps FWHM depending on the energy range windowing (mainly due to intrinsic resolution of the large detector)
Notes: 43

## Slide 44: Pulse Shape Discrimination
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Very good separation between events in LaBr3 and NaI
Events in LaBr3
Events in NaI
Notes: 44

## Slide 45: Energy vs PSD scatter plot
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Notes: 45

## Slide 46
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
SOFTWARE TOOLS AND LIBRARIES
Notes: 46

## Slide 47: Software for the digitizers
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
CAEN provides a wide range of software tools to configure and to control the data acquisition
  Linux and Windows OS compatible (32 and 64 bits)
  Free download
DRIVERS for the communication channel:
  CONET2 – optical link
  VME bus
  USB
LIBRARIES:
  C and LabView
  Demo and examples available for developers
READOUT SOFTWARE:
  Medium and high level tools to manage the configuration and the data acquisition
Notes: 47

## Slide 48: Libraries and drivers
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
A3818
driver
A2818
driver
VME
PCI
PCIe
USB
driver
V1718
driver
USB
CAENcomm library
CONET2 (Optical Link)
CAENDigitizer library
User Applications
USB
Set/Get Params, Start/Stop Read Events, etc.
Open/Close, Read, Write
Notes: 48

## Slide 49: DPP libraries for spectroscopy software
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
Drivers
CAENcomm library
CAENDigitizer library
DPPcore
DPP library
Spectroscopy Applications
HISTO
Digitizers
LIST
WAVE
Acquisition
Server
Hardware
USB, PCI, PCIe
Socket
GUI
Config
Output
Notes: 49

## Slide 50: Next: DPP Event Oriented DAQ
Reproduction, transfer, distribution of part or all of the contents in this document in any form without prior written permission of CAEN S.p.A. is prohibited
x724
x720
x730
x751
…
Readout
SETTING PANEL
Config
Params
Configurator
HW
Find
Devices
Pre-Processor
Waves
Event Selector
Lists
Spectra
RATE MONITOR
SIGNAL
SCOPE
Waveforms
Throttle
Selection Criteria
Histogrammer
E-histo
T-histo
PSD-histo
MCS
Biparam-histo
Readout Server
Data Analysis
GUI
Data Flow
Configuration
DataFile
Time,
Energy,
PSD
Hardware
HISTO PLOTS &
STATISTICS MONITOR
Notes: 50